Common Problems and Their Solutions

| Problem | Probable Cause | Solution |
| --- | --- | --- |
| Sore or dry eyes | Mask not positioned properly. -Mask size not appropriate. | Reapply mask & readjust headgear. -Come in to obtain proper mask size. |
| Redness on the face where mask contacts the skin | Mask on too tight! -irritation or allergy to mask material. | Loosen headgear -Use barrier between skin & mask, such as 3M’s Microspore tape or moleskin tape. -Contact us or physician if problem persists. |
| Runny nose | Nasal reaction to air flow. | If problem persists beyond first 2 weeks of continued CPAP use, contact your physician. |
| Cold nose | Room air temperature is too cold. Air cools while traveling through the tubing | Reposition the tubing so that it runs under your bed covers to reduce heat loss. |
| Dryness or burning sensation in throat or nose | Air is too dry. The relative humidity is less than 40%. | If problem persists beyond first 2 weeks of continued CPAP use: Increase the room humidity. Consider attached humidifier use. Ask your physician for nasal sprays/gels. Make sure you are sleeping with your mouth closed; use a chin strap. |
| Nasal, sinus or ear pain | Sinus infection or middle ear infection. | Stop using CPAP and contact your physician. |
| Feeling of discomfort from the sensation of too much pressure | CPAP pressure above 12.5 cm H20 may cause feelings of discomfort but are necessary to relieve sleep apnea. | It may take up to 4 weeks to adjust to the system. Try to relax when using CPAP. Remember to use the ramp feature. Contact us or physician if problem persists. |
| Pressure delivered seems significantly lower or higher than usual | Possible unit malfunction. (Also possible: adaptation effect; unit under bed w/limited airflow) | Come in to check pressure/unit. |
| Air from CPAP unit seems warm | The filter may be obstructed or dirty. Filter or intake may be occluded by bedclothes, curtains, etc., or room temperature too warm. | Check/replace filter. Move CPAP unit. Turn down thermostat at night. Place unit on floor where unit may be cooler. |